Commodity Description: Fast Protein Liquid Chromatography (FPLC)

| Item No. | Description | System | Quantity |
| --- | --- | --- | --- |
|  | Fast Protein Liquid Chromatography (FPLC) | 1 | 1 |

1. Enduser's Uses

- Fast Protein Liquid Chromatograph (FPLC) for quantitative and preparatory scale purification and polishing of wide range of polysaccharides, proteins, and peptides.

- FPLC system should include necessary consumables for immediate operation and should come with a minimum one-year warranty.

- Supplier should provide installation and the operation/maintenance training to the  user

- FPLC user should be able to operate multi wavelengths of upto 4 in the range of (190-800nm) simultaneously for development of automated purification protocols.

- FPLC  should be able to function during the different chromatography runs with in a pressure range of  0-10MPa

- FPLC should be able to support for screening tests to utilize various  particle size resins   screening for higher resolution and faster purification process.

- FPLC should be equipped with a UV detector with the lowest noise

- FPLC system should be compatible with third party detectors, such as the MALs, particle size analyzer & RI detector.

- FPLC should be equipped with air sensor at appropriate inlet pumps and lines to detect air from buffers and samples.

- FPLC system should be equipped with Fraction collector, binary gradient pump with internally, equipped mixer.

- FPLC should be capable of supporting high throughput screening and other scouting runs.

- FPLC should be capable to handle continuous purification protocols and multicolumn use compatible with different manufacturers.

- FPLC system should have an operating software with following capability such as operation of auto and manual mode of purification protocols methods, cleaning, system and maintenance alarms.

- FPLC should be able to use different method editor options to choose different  Purification templates flow & automation, for  on-line trending.

- FPLC should be able to support online analysis data storage & retrieval.

2. Configuration of goods  (Standard Chromatography Capability/Functions )

2A. System Pumps

- High precision pumps (pump A & B), which delivers buffers in purification runs

- Flow rates ranging from 0.25mL to 10mL/min at operating pressure(0-10Mpa)

- Binary gradient pumps with compatibility to chemical and electric inert

- Dual pumps for small and preparative scale purifications

- Suitable for scale up function

2B. Sample Pump

Automated sample application function from 1mL to 500mL

- Column switching valve

- Suitable for sample loop(1-5mL) and super loop upto (50mL)

- Pressure sensor to read system pressure and alarm signals

- Ease of operation of flow rates from 0.5mL to 5mL/min

- Accuracy of flow rates ± 1-2%

2C. Mixer

- Homogenizes buffers from two system pumps(A&B)

- Should be able to support gradient runs

- Should accommodate varying volumes mixing from (1mL to 25mL) column sizes

- Efficient gradient mixing at 0.5mL to 10mL/min flow rates during gradient runs

- Gradient mixing accuracy of 0.5 to 1%

- A motor to mix with an integrated pressure sensor.

2D. UV/Visible Detectors

- UV/Visible detector to measure wavelength range of 190-800nm

- UV/Visible detector for simultaneous detection of upto four wave lengths.

- Detection of absorbance from 0-3000 AU.

- Status display of absorbance, pH and conductivity values on the screen

- Flow cells with different diameter for preparatory and analytical use (2-10mm)

2E. Conductivity montitor

- Sense buffers and cleaning solutions Conductivity during the purification and clean in place

- Conductivity range of (0-900mS/cm)

- Accuracy of 1± 2%

2F. pH Monitor

- pH monitors for inline sensing of buffers and cleaning solutions pH range of (0-14)

- Easy of pH Calibration

- pH probe for accurate measurement of pH (± 0.1 pH units).

2G. Air Sensors

- Able to detect air from buffers and sample application ends to system and column

- Able to detect and map the bubbles online status during the chromatography operation

- Able to accommodate upto 8 air sensors

2H. Valves

- Inlet A valve upto 8 inlets

- Inlet B valve upto 8 inlets

- Sample inlet upto 15 inlets

2I.  Pressure sensors

- Standard pressure sensors placement to detect below pumps

- System pump

- Sample pump

- Precolumn

- Postcolumn

2J. Column Switching valves

- Able to hold up to 1-5 columns for quick column washing in forward and reverse flow applications.

- The valve should support a column bypass function, allowing the system to be primed or cleaned without removing the column.

- Able to supports up to 2-3Column Switching Valves for up to 1-5 columns

- Each column switching valve should be able to measure pre-column and delta column pressure for the column(s) that are placed in-line.

2K. Outlet valve

- Able to measures the UV/Vis absorbance

- Able to support upto 12 outlet

- collecting fractions in small to large volumes (0.5mL -100mL).

2L.  Fraction Collector

- Suitable for microplates, 15, 50mL

- Should be able to support automated collection of fractions

- Able to support use of multitray racks comprising of different volume holder

3. FPLC system specification

| Parameter | Data |
| --- | --- |
| System configuration | Benchtop system, external computer |
| Control system | Software supported |
| Connection between PC and instrument | Ethernet |
| Dimensions (Length x Depth x Height) | less than 50 x 65 x 80 cm |
| Weight (excluding computer) | 50-100 kg |
| Power supply | 100-240 VAC, 50-60 Hz |
| Power consumption | 100 to 120/220 to 240 V AC or 750 wats maximum |
| Operating temperature range | 4°C to 40°C |
| Pump type | Piston pump, metering type |
| Flow rate range | 0.01 to 100 ml/min (normal range) 0.01 to 50 ml/min (column packing flow) |
| Flow rate specifications | • Accuracy: ±2% • Precision: (Conditions: 0.25 to 10 ml/min, < 10 MPa, 0.5 to 10.8 cP) |
| Mixing principle | Chamber with a magnetic stirrer |
| Mixer volume | 0.6 ml, 1.4 ml (default) or 5 ml |
| Number of monitors | 1 |
| Wavelength range | - 190 -800 nm wavelength range - 5 nm, 2 mm and 10 mm flow cells availability as optional accessories |
| Fraction collector Vessel types | - Deep well plates: 96 / 48 / 24 - 3, 5, 8, 15, 50mL tubes & 250 ml bottles - 250 ml bottles |
| Air Sensors | ≥8 |

4. FPLC Accessories

- Inlet and outlet Valve Kit

- Cables and Pressure tubing

- Fraction Collector enclosure

- Column Valve Kit

- Sample Inlet Valve Kit

- Sample pump

- Shipping Charge

5. Software & PC

- Chromatography software for liquid chromatography should be able to support multiple display traces during evaluation such as absorbance, pH and conductivity.

- Managing chromatography data with more efficient graphical visualization of results.

- Enables calculation of concentration and plot different parameters such as peak integration, pH, conductivity and gradients% etc.

- Chromatogram comparison and peak analysis based on fractions, pH and conductivity and %B

- Multiple user (1-3) software licenses for analyzing the data

- CPU:i5 -4th gen, Intel Core i7 (Quad-Core)

- OS:Windows 10pro (64-bit)

- RAM 16GB, HDP 250GB SSD+1TB HDD

6. Warranty, Installation and training

- 2 years of warranty service to be provided.

- On-site installation by certificated engineer

- Training for operation & maintenance/troubleshooting in English and Korean

7.  Requirements

- IQ and OQ

- Temperature control for fraction collector

- Remote controlling

- Notifications alerts (System function) to user

- Safety certifications

8. ‘위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할수 있음’.