Particle Size analyzer (DLS) Specification

| Item No. | Description | System | Quantity |
| --- | --- | --- | --- |
|  | Particle Size analyzer (DLS) | 1 | 1 |

1. Enduser's Uses

- Particle size analyzer (DLS) for size analysis suitable for polysaccharide, protein and conjugate vaccines.

- Supplier should provide installation and operation/maintenance training to the user

- Supplier should provide IQ/OQ about the instrument.

- Particle size analyzer (DLS) should be able to measure particle sizer range of 0.3 nm - 10 μm

- Particle size analyzer (DLS) should be able to use sample volume under 25 μL for size measurement.

- Particle size analyzer (DLS) should provide three measurement angles (front, side, back)

- Particle size analyzer (DLS) should provide 10mW laser or higher.

- Particle size analyzer (DLS) should provide Multi-angle dynamic light scattering function.

- Particle size analyzer (DLS) should provide auto focusing for measurement

- Particle size analyzer (DLS) should provide able to use sample volume under 25 μL for zeta potential measurement.

- Particle size analyzer (DLS) should provide high conductivity sample measurement for zeta potential.

- Particle size analyzer (DLS) should provide particle concentration measurement function

- Particle size analyzer (DLS) should provide software for all users (multi user software)

2. Particle size analyzer (DLS) system specification

| Parameter | Data |
| --- | --- |
| System configuration | Benchtop system, external computer |
| Control system | Software supported |
| Connection between PC and instrument | USB or another port |
| Dimensions (Length x Depth x Height) | less than 80 x 60 x 40cm |
| Weight (excluding computer) | ~30 kg |
| Power supply | 100-240 VAC, 50-60 Hz |
| Temperature control range | 0°C to 80°C |
| Laser | ≥10mW |
| Detection angle | Front, Side, Back |
| Detection angle number for measurement | 1 and 3 (Multi angle dynamic light scattering) |
| Minimum sample loading volume for measurement | ≤25 μL |
| Focusing for measurement | Auto focusing |
| Size measurement range | 0.3 nm - 10 μm |
| Maximum sample conductivity for zeta potential measurement | 250mS/cm |
| Particle concentration measurement range | ~1x1012 or higher |

3. Software & PC

- Software for Particle size analyzer (DLS) should be able to support operating and evaluating results.

- Managing data with more efficient graphical visualization of results.

- Multiple user (~8) software licenses for analyzing the data.

- CPU: 4th gen+, Intel Core i7 or equivalent

- OS: Windows 10pro (64-bit)  or equivalent

- RAM 16GB, 1TB HDD

- Monitor

4. Warranty, Installation and training

- 1 year of warranty service to be provided.

- On-site installation by certificated engineer

- Training for operation & maintenance/troubleshooting in English and Korean

5.  Requirements

- IQ and OQ

6. ‘위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할수 있음’.