Commodity Description

| Item No. | Product Description | Unit | Q’ty |
| --- | --- | --- | --- |
| 1 | NSF-Certified Type A2 Biosafety Cabinet | Set | 9 |

The Type A2 Biosafety Cabinet is a state-of-the-art containment solution designed to provide both product protection and environmental safety. Certified by the National Sanitation Foundation (NSF), this cabinet meets rigorous standards for personnel, product, and environmental protection, making it ideal for handling microbiological and biomedical materials where the user and the work environment need protection against biological hazards.

Key Features

- Certification: NSF-certified, ensuring compliance with strict safety and performance standards.

- Airflow System: Vertical laminar airflow with a minimum inflow velocity of 100 fpm (feet per minute) and downflow velocity of 55 fpm, maintaining a stable and contamination-free environment.

- ULPA Filtration: Equipped with Ultra Low Penetration Air (ULPA) filters capable of trapping particles as small as 0.12 microns with 99.999% efficiency, surpassing traditional HEPA filters in both particle retention and cleanliness levels. Construction: Constructed from stainless steel with a corrosion-resistant powder-coat finish, offering durability and ease of cleaning.

- Safety Features: Includes an audible and visual alarm to alert users to any airflow disruptions, ensuring maximum safety during operation.

- Energy Efficiency: Features an energy-efficient motor to reduce operational costs and environmental impact.

Specifications

- Noise Level: Less than 70 dB

- Power Supply: 220 V

- Size: 4 ft

Additional Components

- UV Light: UV lamp for decontamination processes.

- Service Valves: Built-in service valves for gas, air, and vacuum services to accommodate a wide range of research needs.

- Utility Ports: Multiple built-in ports for easy integration of additional equipment.

Warranty and Support

- Warranty: 5-year warranty on all components, with extended service options available.

Delivery and Installation

- Lead Time: Within 3 months from the contract date.

- Installation: Comprehensive on-site installation by certified technicians, including operational training and certification for use.

- Equipment Disposal: For the purchase of 9 units, 5 will be provided under the condition that the existing equipment being replaced is disposed of and removed at the time of new equipment installation.

- Validation: All biosafety cabinets will include complete validation services to ensure full operational compliance and performance after delivery.

위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할 수 있음.