COMMODITY DESCRIPTION

| HSK NO. | ITEM NO. | DESCRIPTION | UNIT | QUANTITY |
| --- | --- | --- | --- | --- |
|  |  | -80℃ Ultra-Low Temperature Freezer (Independent Cooling System, 독립냉각) |  | Total 15 |

Ⅰ. INDEPENDENT cooling system – 15 SETS

- Features

- Utilizing two independent cooling structures, this cooling system maintains an internal temperature of -80℃ even in the event of failure of one compressor.

- Minimize the operational stress of the system to enhance its exceptional reliability and ensure sample protection.

- The inner door prevents loss by efficiently sealing cold air through an effective mechanism.

- Through the emission of minimal heat into the environment, it optimizes energy efficiency by reducing the workload on cooling, heating, and air conditioning systems.

- Samples remain securely protected within the cabinet, with temperature differentials maintained below 5℃.

- Operating quietly with noise emission levels below 60 dBA, the freezer maintains a stable environment in hallways.

- It provides systematic storage methods utilizing standardized storage accessories such as racks, trays, and boxes and should be included in the estimate.

- Search for the current temperature status of the freezer and the record of past operational temperature data on the screen.

- Through a mobile app, users can check and control the status of the ultra-low temperature freezer anytime, anywhere, using devices such as computers and laptops.

- Providing access to crucial data and equipment alerts, including equipment and ambient temperatures, alarms, and warning notifications.

- By utilizing the location tracking and memo features for samples stored in the ultra-low temperature freezer, the frequency and duration of door openings are reduced.

- Security is maintained through the use of biometric recognition, card key systems, and keys.

- The freezer should be capable of accommodating the installation of CO2 and LN2 backup systems at a later.

- Over the past five years, there has been a track record of supplying equipment with data transmission systems.

- Specifications

- Temperature Range: -60℃to -90℃

- Vial Capacity: 40,000 vials or more (The quantity of vials based on a 9*9 box)

- Weight (excluding rack): 450 kg or less

- Cabinet type/ Door type: Upright/ Single

- Electrical data: 110V, 220V 60Hz, 230V 50Hz

- Remarks

- The installation and operation should be carried out by the bidder/supplier responsibility at end-user's site.

- IQ and OQ must include.

- More than two years warranty service should be provided by supplier after the performance test.

- 위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할 수 있음.