Specification for ARF IVC

(Single-sided Individually Ventilated Cage System)

- General Specifications

- Description: Individually Ventilated Cage system for Mice

- Component: [2 Single sided rack with one Air blower] * 2 sets including 50% of spare cage sets.

- Dimension(mm): less than 1800 W * 550 D * 2200 H / rack

- Cage capacity: total 60~70 cages / rack

- Cage set: cage body, wire bar lid, filter-top cover, cage card holder (water bottle and nozzle excluded).

All materials including cage set could be autoclavable.

- Rack Frame, wire lid material: STS 304

- Cage material (Cage body, filter-top cover): clear Polysulfone

- HEPA grade filter should be inserted into cage top.

- The air is supplied through the pre-filter and HEPA filter as well as exhaust.

- The exhaust air always goes through the connection for reducing the odor.

- The supply/exhaust nozzles have heat resistance and chemical resistance, and make the animals reduce stress and minimize the dirt.

- In case both of the cages are on the rack and the removal from the rack, containment is guaranteed.

- External water bottle replaceable without removing the cage from the rack reduce cage opening frequency.

- Every part except blower should be easily disassembled and assembled for washing, autoclaving or chemical disinfection.

- Performance Specifications

- Standard operation of Air Change per hour (ACH): more than 50/hr

- Efficiency of HEPA filter of Air Handling Unit: more than 99.97%, 0.3 ㎛

- Multi linking capability: one AHU serving up to 2 single-sided racks.

- Chemical resistant and autoclavable.

- Temperature and humidity sensor should be installed inside the ventilation unit.

The Temperature and Humidity all the time monitoring on the screen.

- Replacing the HEPA filters monitoring

- Noise level, vibration of blower should be minimized to animal < 55 dBA.

- Positive or negative pressure inside the cage should be achieved over a certain level.

Display the pressure, ACH in each cage all the time monitoring on the screen.

- The system should alarm in event that one of temperature, humidity, pressure, or ACH in the cage.

- The IVC RACK Airflow should be controlled stably regardless of inserted cage quantity. Also, pressure should maintain either positive pressure or negative pressure over certain level even though more than 90% of cages are removed from the racks.

- Remarks

- Three years warranty service will be provided.

- Training is provided for the users.

- The installation and operation test must be performed at IVI by supplier.

- Acceptance Test Run must be confirmed by final user at IVI.

- Concise user instruction shall accompany all units in the form of a user manual.

- Validation reports made by internationally accredited organizations should be provided. (Pressure, Dead zone, Velocity and Noise etc.)

- HEPA Filter installation certification should be provided.

- Delivery date: within 4 months from the date of contract

- 위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할 수 있음.