Commodity Description

| Item No. | Product Description | Unit | Q’ty |
| --- | --- | --- | --- |
| 1 | CO2 incubator with water tray | Set | 2 |

A.  Description

CO2 incubators are designed to provide a stable as well as homogeneous environment needed for mammalian cell cultures by controlling temperature and carbon dioxide to facilitate optimal cultivation conditions.

A water tray included incubator is needed to eliminate the risk of contamination and to ease the clean the chamber.

Moreover, it should have fast recovery of temperature and CO2 gas within 10minutes or less. It includes audible and visual alarms to notify the user in case of problems such as water or CO2 shortage, ensuring maximum safety during operation.

B. Specifications

- High temperature disinfection: 180℃ or equivalent

- Capacity: 150L ~ 170L

- Exterior Dimensions (WxDxH): Within 70 x 70 x 90 cm

- Net weight (kg): Within 105 kg

- Temperature range: 4℃ above ambient to 50℃

- Temperature stability at 37℃: ±0.1℃

- CO2 range: 0.1-20% or equivalent

- CO2 stability at 5% CO2 : ±0.1%

- CO2 Recovery: 10minutes or less

C. Accessories

- Water tray

- Shelf

D. Remarks

- The water tray must be separated.

- The handle should be on the right side.

- A warranty service of two years should be offered.

- Training for the users should be provided.

- The supplier must carry out the IQ/OQ service at IVI.

- The final user at IVI must confirm the Acceptance Test Run.

- A concise user manual should be included with all units.

- Validation reports from globally recognized organizations should be supplied.

- A certification for the installation of the HEPA Filter should be provided.

- Delivery date: within 1 month from the date of contract

E. 위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할 수 있음.