COMMODITY DESCRIPTION

| HSK NO. | MODEL NO. | DESCRIPTION | UNIT | QUANTITY |
| --- | --- | --- | --- | --- |
|  |  | CO2 incubator | EA | 6 |

A. Features

- Fast temperature and CO2 recovery in less than 10 min after door opening, with no setpoint overshoot

- Easy-to-clean, stainless-steel chamber without seams, fan, ducts, covers, etc.

- 180°C sterilization

- Comprehensive documentation of CO2, temperature, door opening times, alarms, etc. for easy troubleshooting in case of alarms

- GMP compatible: ISO class 4 / GMP grade A clean room compatible

- Readily stackable in a compact footprint

B. Specifications

- Chamber volume: 165L or greater

- External dimensions (W*H*L): 718*715*900 mm or less

- Shipping weight: 113 kg or less

- Max. shelf load: 10 kg

- Sterilization temperature/duration: 180°C / under 12 hours or similar

- Temperature control: ±0.1°C

- Temperature uniformity: ±0.3°C

- CO2 range/control: 1 – 20%/±0.1%

- CO2 sensor: T/C sensor or equivalent

- Interior material: Type 304, polished stainless steel or equivalent

- Power supply: 220V, 50/60Hz

C. Remarks

- Installation and operation are carried out directly by the supplier at the installation site.

- User manual (English or Korean) is provided.

- 2 years of warranty after the performance test (excluded consumable parts).