Commodity Description: High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) dual pump system

| Item No. | Description | System | Quantity |
| --- | --- | --- | --- |
|  | High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) dual pump system | 1 | 1 |

1. Enduser's Uses

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) for size analysis suitable for polysaccharide, monosaccharide and conjugate vaccines.

- Supplier should provide installation and operation/maintenance training to the user

- Supplier should provide IQ/OQ about the instrument.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should be provide dual quaternary pump system.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should be provide ~10mL/min flowrate.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should be provide over 35MPa working pressure.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should be provide resolution of PAD detector under 0.0025nS/cm.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide maximum detector cell pressure 10MPa.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide maximum detector cell temperature 60℃.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide able to use sample volume under 200 μL for analysis.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide autosampler temperatures control function (4℃~50℃).

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide Dual quaternary pump system.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide RI detector range -0.004~+0.004.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide RI detector dynamic range over 10000000:1.

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should provide software for all users (multi user software).

2. High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) dual pump system specification

| Parameter | Data |
| --- | --- |
| System configuration | Benchtop system, external computer |
| Control system | Software supported |
| Connection between PC and instrument | USB or another port |
| Power supply | 100-240 VAC, 50-60 Hz |
| Temperature control range of autosampler | 4°C to 50°C |
| Minimum sample loading volume for measurement | ≤200 μL |
| Pump type | Dual quaternary pump |
| Pump working pressure range | 0~35MPa |
| Pump flow rate range | 0~10mL/min with standard PEEK pump head |
| PAD detector resolution | ≤0.0025nS/cm |
| PAD detector sample rate | 1~100 Hz |
| RI detector range | -0.004~+0.004 |
| RI detector dynamic range | ≥10000000:1 |

3. Software & PC

- High-Performance Anion-Exchange Chromatography with Pulsed Amperometric Detection (HPAEC-PAD) should be able to support operating and evaluating results.

- Managing data with more efficient graphical visualization of results.

- Multiple user software licenses for analyzing the data.

- CPU: 4th gen+, Intel Core i7 or equivalent

- OS: Windows 10pro (64-bit)  or equivalent

- RAM 16GB, 1TB HDD

- Monitor

4. Warranty, Installation and training

- 2 years of warranty service to be provided.

- On-site installation by certificated engineer

- Training for operation & maintenance/troubleshooting in English and Korean

5.  Requirements

- IQ and OQ

6. ‘위 규격서에서 명시한 제품과 동등 이상의 다른 제품 납품이 가능한 업체는 입찰에 참가할수 있음’.