Commodity Description: Multi Angle Light Scattering (MALS) System

| Item No. | Description | System | Quantity |
| --- | --- | --- | --- |
|  | Multi Angle Light Scattering (MALS) System | 1 | 1 |

1. Enduser's Uses

- Multi Angle Light Scattering (MALS) System for molecular weight and size analysis suitable for polysaccharide, protein and conjugate vaccines.

- Supplier should provide installation and operation/maintenance training to the user

- Supplier should provide IQ/OQ about the instrument.

- Multi Angle Light Scattering (MALS) System should provide 120mW GaAs linearly polarized laser.

- Multi Angle Light Scattering (MALS) System should provide over 10,000 laser lifetimes.

- Multi Angle Light Scattering (MALS) System should provide at least 18 detector angles.

- Multi Angle Light Scattering (MALS) System should provide 24-bit detector resolution.

- Multi Angle Light Scattering (MALS) System should provide self-cell cleaning system.

- Multi Angle Light Scattering (MALS) System should provide molecular weight range up to 1G g/mole (daltons).

- Multi Angle Light Scattering (MALS) System should provide molecular size range 10 to 500nm and up to 1000nm with shape-specific models.

- Multi Angle Light Scattering (MALS) System should provide sensitivity to 10ng of polystyrene and 200ng of BSA.

- Multi Angle Light Scattering (MALS) System should provide compatibility with Optilab RI detector.

- Multi Angle Light Scattering (MALS) System should provide ASTRA software.

2. Multi Angle Light Scattering (MALS) System specification

| Parameter | Data |
| --- | --- |
| System configuration | Benchtop system |
| Control system | Software supported |
| Connection between PC and instrument | USB or another port |
| Power supply | 90-250V @ 50-60Hz, AC |
| Light source | 120mW GaAs linearly polarized laser |
| Laser life time | ≥ 10,000 hours |
| Detectors | ≥18 angles |
| Detector resolution | 24 bit |
| Mobile phase compatibility | All-solvent compatibles (aqueous and organic) |
| Molecular weight range | ≤ 200 ~ 1G g/mole (daltons) |
| Molecular size range | 10 ~ 500nm, up to 1000nm with shape-specific models |
| Sensitivity | 10ng of 100kDa polystyrene, 200ng of BSA |

3. Software

- Managing data with more efficient graphical visualization of results.

- Multiple user software licenses for analyzing the data.

- ASTRA software

4. Warranty, Installation and training

- 1 year of warranty service to be provided.

- On-site installation by certificated engineer

- Training for operation & maintenance/troubleshooting in English and Korean

5. Requirements

- IQ and OQ

- Colum set for analysis