Commodity Description: Ultra High Performance Liquid Chromatograph

| Item No. | Description | System | Quantity |
| --- | --- | --- | --- |
|  | Ultra High Performance Liquid Chromatograph | 1 | 1 |

1. Enduser's Uses

- Ultra High Performance Liquid Chromatograph for analysis suitable for polysaccharide, protein and conjugate vaccines.

- Supplier should provide installation and operation/maintenance training to the user

- Supplier should provide IQ/OQ about the instrument.

- Ultra High Performance Liquid Chromatograph should provide quaternary pump.

- Ultra High Performance Liquid Chromatograph should provide maximum pressure range 15,000 psi with a flow rate above 5mL/min.

- Ultra High Performance Liquid Chromatograph should provide flow accuracy ±0.1%.

- Ultra High Performance Liquid Chromatograph should provide Low-pressure gradient proportioning.

- Ultra High Performance Liquid Chromatograph should provide sampler maximum pressure range 15,000 psi.

- Ultra High Performance Liquid Chromatograph should provide injection volume accuracy  ±0.5% for 10 µL water.

- Ultra High Performance Liquid Chromatograph should provide sample compartment range 4~40℃.

- Ultra High Performance Liquid Chromatograph should provide Diode array detector(DAD).

- Ultra High Performance Liquid Chromatograph should provide DAD wave range 190~800nm.

- Ultra High Performance Liquid Chromatograph should provide DAD wavelength accuracy ±1nm.

- Ultra High Performance Liquid Chromatograph should provide DAD wavelength repeatability ±0.1nm.

- Ultra High Performance Liquid Chromatograph should provide a function to connect and replace tubing within the instrument without tools.

- Ultra High Performance Liquid Chromatograph should provide connection compatibility with WYATT MALS (DAWN) system.

- Ultra High Performance Liquid Chromatograph should provide Chromeleon software for operating and data analysis.

2. Ultra High Performance Liquid Chromatograph specification

| Parameter | Data |
| --- | --- |
| System configuration | Benchtop system |
| Control system | Software supported |
| Connection between PC and instrument | USB or another port |
| Pump | Quaternary pump |
| Chromatography system maximum pressure | ≥15,000 psi |
| Flow accuracy | ±0.1%. |
| Gradient Formation | Low-pressure gradient proportioning |
| Sampler maximum pressure | ≥15,000 psi |
| Injection volume accuracy | ±0.5% for 10 µL water. |
| Detector | Diode array detector (DAD) |
| DAD wave range | 190~800 nm |
| DAD wavelength accuracy | ±1nm |
| DAD wavelength repeatability | ±0.1nm |

3. Software

- Managing data with more efficient graphical visualization of results.

- Multiple user software licenses for analyzing the data.

- Chromeleon software

4. Warranty, Installation and training

- 1 year of warranty service to be provided.

- On-site installation by certificated engineer

- Training for operation & maintenance/troubleshooting in English and Korean

5. Requirements

- IQ and OQ

- TSKGel column 2 set (column set: G3000PW, G5000PW, G6000PW, Guard column)